# Flexible Conductive Silicone Pressure Sensor Fabrication Protocol

Purpose
This document describes the fabrication process used to create a flexible conductive silicone pressure sensor consisting of a conductive polymer composite sheath surrounding silicone tubing and encapsulated within an outer silicone layer.

## Materials

Silicone Rubber Part A (Let’s Resin, Hardness 0A)

Silicone Rubber Part B (Let’s Resin, Hardness 0A)

Precision Coated and Converted Fibers (PC2F / NCCF), Product #1AC1PC-0.5

Branched Nickel Strands (BNS), Product #3AA150

Silicone Oil (MicroLubrol Type 200 Fluid, 50 cSt)

Silicone tubing (approximately 3 mm ID, Walmart source)

Trigger clamps

Stainless steel rod/needle stock sized to fit through tubing lumen

Scale

Vacuum chamber

Beakers

Glass stirring rod

Disposable weighing boats

Pipettes (tips cut to act as scoops)

Scissors

3D printed molds

## 3D Printing Information

Printer: Original Prusa XL – 5T Input Shaper, 0.4 mm nozzle

Filament: Prusa PLA

Include STL files directly in the GitHub repository.

## Composite Formulations

| Recipe | Part A | Part B | NCCF | BNS | Silicone Oil |
| --- | --- | --- | --- | --- | --- |
| 1 | 35.6% | 35.6% | 4.7% | 19.6% | 4.5% |
| 2 | 39.7% | 39.7% | 13.2% | 4.1% | 3.3% |

Typical batch size: 4 g Silicone Part A and 4 g Silicone Part B. Fifty percent of Part A and Part B are reserved for fabrication of the outer silicone encapsulation layer.

## Mold Files

Store all STL files in the GitHub repository. Suggested structure: Conductive Composite Top Mold, Conductive Composite Bottom Mold, Silicone Top Mold, Silicone Bottom Mold.

## Fabrication Procedure

Measure all materials by weight using the selected formulation.

Cut pipette tips to use as scoops for dry materials.

Mix silicone oil and BNS thoroughly, then divide equally into two beakers.

Divide NCCF into two equal portions. Add one portion to Silicone Part A and the other to Silicone Part B.

Combine one silicone mixture with one BNS/oil mixture. Repeat for the second beaker.

Mix both beakers until homogeneous.

Combine the contents of both beakers and stir for 15 minutes.

Place the mixture in a vacuum chamber. Apply sufficient vacuum to remove trapped bubbles (approximately 25 Pa indicated on the chamber gauge) for 30 minutes.

While degassing, cut a 2.5 inch section of silicone tubing and insert a stainless steel rod through the lumen to maintain straightness.

Fill the conductive composite bottom mold with the degassed composite.

Place the tubing into the mold slots.

Fill the top mold with composite material and assemble the mold.

Clamp the mold using trigger clamps.

Cure at room temperature for at least 4 hours.

Mix equal parts Silicone Part A and Silicone Part B for the outer encapsulation layer.

Mix for approximately 1 minute and degas for 30 minutes using the same vacuum process.

Remove the cured conductive composite sheath from the mold and trim excess material.

Fill the silicone bottom mold with the degassed silicone.

Place the conductive composite sheath into the mold.

Fill the silicone top mold and assemble the mold.

Clamp the mold and add additional silicone through the fill holes until completely full.

Cure at room temperature for at least 6 hours.

Demold the sensor and trim excess silicone.

## Acceptance Criteria

No visible air bubbles or large voids.

Tubing lumen remains open and unobstructed.

Conductive sheath is continuous around the tubing.

Outer silicone layer fully encapsulates the conductive composite.

Sensor can be removed from the mold without tearing.